NEW ALTERNATIVES FOR PATIENTS SUFFERING IN-HOSPITAL CARDIAC ARREST: THE ROLE OF FULLY AUTOMATIC DEFIBRILLATORS

A. Martinez-Rubio1 , N Kanaan2, M Borggrefe3, M Block4, M Makijarvi5,

F Fedele6, C Pappone7, W Haverkamp8, JL Merino9, J Cinca10

1Hosp. de Sabadell (Parc Tauli), Barcelona, Spain, 2Cardiac Science, Irvine, California, USA, 3Univ Hosp of Mannheim, Mannheim, Germany, 4Stiftsklinik Augustinum, Muenchen, Germany, 5Univ Hosp of Helsinki, Helsinki, Finland, 6Univ Hosp of Rome, Rome, Italy, 7Hosp San Raffaele, Milan, Italy, 8Univ Hosp of Muenster, Muenster, Germany, 9Univ Hosp La Paz, Madrid, Spain,

10Univ Hosp of St. Pau, Barcelona, Spain

Introduction: The response time to cardiac arrest is a major determinant of morbidity and mortality. Methods: The safety and accuracy of the response of a new fully automatic external cardioverter defibrillator was examined in an European multicenter trial by reviewing the simultaneously recorded Holter data and the programmed parameters. Patients at risk of cardiac arrest in monitored wards (n=51) and patients undergoing electrophysiological testing or ICD implantation (n=66) were enrolled.  Results: During 1,240 hours, 1,988 episodes of rhythm changes were documented. The device detected ventricular tachyarrhythmias with a sensitivity of 100% and specificity of 97.6% (true negatives n=1454, true positives n=499, false positives n=35, false negatives n=0). The false positives were all caused by T-wave oversensing during ventricular pacing.  There were no complications or adverse events. The mean response time was 14.4 seconds for those episodes needing a fully charge of the capacitor. Conclusions: The use of this safe and fully automatic device will result in a short response time to treatment and it is likely that it will significantly improve the outcome of patients with in-hospital cardiac arrest.