MODIFIED ALTERNATING CURRENT DEFIBRILLATION-A NEW HOME-DEFIBRILLATION TECHNIQUE FOR SUDDEN DEATH

A.T. Weiss, I. Katz, S. Gorni, S. Rosenheck

Hadassah University Hospital, Jerusalem, IL, USA

Background: Defibrillation is the only clinically effective treatment of ventricular fibrillation (VF). Only prompt defibrillation improves outcome and increases the chance of survival with full recovery. Immediate availability of a home-based defibrillator using mains-derived Alternating Current (AC) current will drastically improve outcome. The aim of our study was to develop a defibrillator based on modulated AC current, resembling biphasic configuration and compare its efficacy, in a pig model to a standard Direct Current (DC) Defibrillator.

Methods: A computer controlled, modulated AC defibrillation system was developed using a High Voltage Switch and a High Voltage Transformer. The efficacy and safety was evaluated in 5 pigs (30-40 Kg), under general anesthesia with ketamine and isoflouran. A single quadripolar-pacing catheter was inserted percutaneously. VF was induced with rapid ventricular burst pacing and stable VF was defibrillated after 15 seconds. Defibrillation threshold (DFT) was determined in each animal with AC and standard DC shock using step-down protocol.

Results: A biphasic like shock was used with a short isoelectric stage between the phases. The DFT with AC was 70.83±24.81 Joules and with DC was 65.83±12.41 Joules (p=0.49). No macroscopic damage was observed after AC or DC defibrillation.

Conclusions: Modulated AC defibrillation is safe and effective as the commercially available DC defibrillation. The defibrillator is built from inexpensive high voltage transformer, without need for capacitor, batteries or routine maintenance, and delivers repeated shocks without any delay. Therefore it may be an ideal platform for automatic home defibrillator to improve outcome of sudden death.