LONG-TERM EVALUATION OF BIOTRONIK LINOX AND LINOXSMART ICD LEADS

E.D. Good6, I. Cakulev1, M. Orlov2, D. Hirsh3, H. Bloom4, J. Simeles5, K. Mohr5

1. University Hospitals of Cleveland, Cleveland, OH, USA

2. Caritas St. Elizabeth's Boston, Boston, MA, USA

3. Emory University, Atlanta, GA, USA

4. Emory University and Atlanta VA Medical Center, Atlanta, GA, USA

5. BIOTRONIK, Portland, OR, USA

6. University of Michigan, Ann Arbor, MI, USA

Introduction: Spurred by a rash of recent ICD lead failures and advisories, systematic assessment of post-implant lead performance has risen to the level of imperative in the collective conscience of both physicians and industry, alike.  GALAXY (NCT00836589) and CELESTIAL (NCT00810264) are ongoing multi-center, prospective, non-randomized observational studies being used to effect such evaluation by analysis of long-term safety and performance data of BIOTRONIK leads.

Methods: ICD and CRT-D pts are being followed for lead performance and safety of the Linox and Linoxsmart ICD leads for 5 yrs post-implant. All procedural and system-related adverse events (AEs) are assessed at each follow-up along with lead electrical parameters.  An independent CEC of 5 EPs adjudicates all AEs to determine AE category and lead relatedness. Adjudicated AEs are then categorized per ISO 5841-2 (3rd edition). Figure 1 displays a Kaplan-Meier actuarial graph by Linoxsmart and Linox ICD model groups.

Results:  A total of 3915 leads were implanted in 3833 pts (73.0% male, mean age 67.0 ± 12.2 yrs) at 146 US centers.  The estimated cumulative survival probability is 97.3% at 5 yrs after implant for Linox leads and 98.3% at 4 yrs after implant for Linoxsmart leads. The most common AEs were oversensing (19, 0.49%), failure to capture (10, 0.26%), and conductor fracture (9, 0.23%), abnormal pacing impedance (7, 0.18%), lead dislodgement (7, 0.18%), insulation breach (5, 0.13%).

Conclusion: Linox and Linoxsmart ICD lead-related AEs are rare; and, lead electrical performance and cumulative survival probability are excellent.