# 93

SINGLE-LEAD VDD ICD: A CASE REPORT

O.T. Greco A. Cardinalli Neto   M.J.F. Soares   C.A. Santos   A.C. Brandi   J.C.M. Brandi   P.R. Pavarino   P.F. Mata   T. Domingos   D.M. Braile

1=Instituto de Molestias Cardiovasculares, Sao Jose do Rio Preto, SP, Brazil          10=

Background: The use of VDD mode in bradytherapies, showed good results. This case report relates the use of a similar, but innovative lead that incorporates a shock coil. A specific ICD was designed to deal with this VDD lead. This system permits an increased specificity on supraventricular tachycardias (ST) discriminatory algorithms, similarly observed in double-chamber ICDs. Case Report: The patient is 72 years old, was admitted with cranial traumatism, due to a syncope followed by a ground fall. The 24h Holter indicates sinus rhythm with CHB and episodes of sustained ventricular tachycardia. The patient received the Deikos A+ with Kainox VDD in Sept/2001. Results: The values measured were: P-wave = 1.6 mV; R-wave = 9.5 mV; atrial threshold = 0.9 V; ventricular threshold = 0.6 V; atrial impedance = 458 W; ventricular impedance = 689 W; defibrillation threshold = 14 J. Conclusion: Preliminary data showed demonstrate the viability of this approach. It’s possible to have the same advantages observed in double-chambers ICDs, related to specific algorithms used to discriminate the ST episodes.

# Text boxes

212